## CMS
Manufacturer: | Mindray | Philips | GE | Draeger | Spacelab | Nihon Kohden | Nihon Kohden
Product Name: | BENEVISION CMS | IntelliVue Information Center iX | CareScape Central Station | Infinity Central Station Wide | Xhibit | CNS-6201 | CNS-9101
Max no./display | 64 | 32 | 16 | 16 | 16, Max resolution supported : | 16 | 24
Max no./central station | 64 per WS / VS and Hardware CMS | 32 | 64 | 64 , 32 in front, 32 in background | 48 | 32 | 48
Max no of beds per virtual CMS | 600 | 1600 | ? | N/A | ? | ? | ?
Type | 4 ordinary displays with 1280×1024 resolution 4 widescreen displays with 1920×1080 resolution Support single 4K widescreen with 3840X2160 resolution (recommended displays: 55 inches or above with aspect ratio of 16:9) Support touchscreen | Flat-panel LCD Touch or not , 1280 x 1024; larger display sizes require higher resolutions (1920 x 1080, 2560 x 1440, 2560 x 1600) | Wide Flat-panel touch or Wide Panel PC touch | 22"Touch Flat-panel LCD | 22"Touch Flat-panel LCD | Flat-panel LCD touch, 24" resol 1920*1200 | Flat-panel LCD, Touch, 24" (resolution 1920x1080)
Single or dual displays | 1, 2, 3 or 4 | Single and dual | Single or dual | Single or dual | 1, 2, 3 or 4 | Single or dual | Single or dual
Size,(in) | 23" standard, can offer larger one | 19"up to 42" | 21" | 22" (1680x1050) | 22" | 24" | 24"
Remote Display and Control | WorkStation 8 / 16 / 32 / 64 beds (exclusive app running on Windows OS) WorkStation Lite (not exclusive app running on Windows OS - TM80 not supported) | Yes, 32 beds | Only Display, no control capabilities | Only Display, no control capabilities | Yes | yes | yes
Non-interactive (slave) | ViewStation, 64 beds can print, review | Yes, 32 beds | Yes via video splitter (Mirror) | Yes via video splitter | Yes via video splitter | Yes | yes
Remote viewing (outside hospital) | Windows App - CMSViewer Browser - WebServer Mobile App - MobileViewer | Yes (web access) | Yes with mobile Viewer suite ( Web Viewer, Pocket and cellular viewer) | Patient Watch (?) | Yes | yes ( net connect server) estimated price : 11000 E +PC to buy | yes ( net connect server)
Mobile App | iOS and Android - MobileViewer Android - AlarmGUARD | iOS and Android | IOS /Android ? | No | ? | IOS | IOS
Multiview web access | Yes (64 beds? - to be released in Q2 2025) | Yes (up to 32 on one browser page) | yes up to 16 on one web window | No | ? | Yes | yes
Pairing solution : more than one equipment per sector | Yes - max no to be checked PM + Vent / Ane + Pumps Defib + Vent / Ane + Pumps | Yes, up to 5 | Yes , only tele paired with monitor | No | ? | ? | ?
Max no. per patient sector | 19 | 50 (configurable) | 16 ? | 12 | All | 8 | 8
Max no. per individual patient display | 26 | 50 (configurable) | 17 ? | 24 | All | All | 56
Max no./patient sector (main display) | 7 | 12 | 4 | 4 | 2, if in 16 patients | 16 | 16
Max no. (individual patient display) | 12 | 32 (all) | 8? | 8 + 12 for 12L ECG | All | 16 | 16
Graphical | Yes | Yes , split trends and ST-MAP within sector | Yes , split trends within sector | Yes | Yes | Yes
Tabular | Yes | Yes | yes | Yes | Yes | Yes
Length of time | 10 days | ≤7 days | 72H (3days) | 5 days max, depending on bedside monitor | 4 days max (depends of the connected monitor) | 5 days
Trend upload from monitor (once reconnected to network) | Yes | Yes (Trend upload , 8H) | Yes, needed a dedicated server | ? | No | yes | yes
Full disclosure upload from transport monitor (once replug to host monitor) | No (Event only) | No | No | No | No | Yes | Yes
ARRHYTHMIA ANALYSIS | 24 arrhythmia alarms | 25 enhanced arrhythmia alarms
Bedside monitors | Yes : HR, ST, PVCS, QT/QTc, RR, SpO , PR, NIBP, TEMP, IBP, CO, EtCO , Multi-gas, O , N O, CCO, ICG, RM, BIS, tcGas, EEG, NMT, rSO2 | Yes | yes | yes | Yes | Yes
Telemetry | Yes | Yes | yes | Yes | Yes | yes
No. of leads analyzed simultaneously | 2 if Mindray 1 if Mortara (UHF Telemetry only) | 2 | 4 | 2 max | 2 for bedside, 3 for Xhibit telem | ≤12, varies
12L ANALYSIS | Yes | yes up to 100/patients | Yes | Yes ( Glasgow algorithm) |  | yes up to 200/patients, kept for 5 days max | yes up to 200/patients, kept for 5 days max
Bedside monitors | Yes | Yes | Yes | yes | Yes | Yes
Telemetry | Yes | Yes | Yes | Yes | Yes | yes
No. of leads analyzed from bedside monitors | ≤12 | 12 | ≤12 | 12 l max | 12 | ≤12, varies
No. of leads analyzed from telemetry | 12 (from 6 leads) | 12 (HEXAD, EASI) | ≤8 | 12 L TruST |  | 8 max | 8
Ventilator parameters ( alarms settings ) + PV loop | No | yes from EC10 or Spiro Module,No PV loop | yes from Unity Network or Spiro Module,No PV loop | Yes: Vent Central app |  | No | No
QRS tone for one selected patient | No | No | No | No | No | Yes | Yes
Bedside monitors/telemetry | Yes/Yes | Yes/Yes | yes/yes | Yes/Yes | Yes/Yes | Yes/yes | Yes/yes
Type | All | Physiologic and events selectable | Arrhythmia, ST, and Sample | Arrhythmia and alarm events, manual | Arrhythmia, ST, OxyCRG | Alarm history, ST recall, hemodynamics, arrhythmia recall, interpretive 12-lead ECG
Max no. of events | 3000 | Unlimited, all events sent from connected devices are stored | 2000/patient | 1000 / patients | Unlimited needs ICS | 10,000 files event history; 7,200 files ST recall; 256 files hemodynamics; 1,500 files arrhythmia recall; 200 files interpretive 12-lead ECG, per patient
Length of time | Unlimited | ≤7 Days | NA; event storage treated on a first-in first-out basis; limits dependent on connected device(s) | 5 days | Unlimited needs ICS | 5 Days
Length of time | 10 Days 256 waves : 4 waves/patients max | ≤7 Days all monitored waves (24 waves /patients) | ≤ 6 days | 5 days | 3 days | 5 Days
No. of parameters stored | All | Unlimited, all parameters sent from connected devices are stored | 9 waves | Patient source depending | All | All
No. of waveforms stored | All | ≤32/patient | ECG12L + Resp/Sp02/IBP | 16 max/patient | All | 16
Diagnose 12L | Yes | Yes | ? No | Yes | Yes | Yes | Yes
OxyCRG waveform | Yes, up to 100 events ( not full disclosure) | No | ? No | No |  | ? | ?
Patient Data Storage | Unlimited Storage | Unlimited Storage |  | N/A
Databases | Proprietary | MS SQL Enterprise
Type | All | High and low for all monitored parameters and arrhythmias | Varies by parameter | Audible and visual |  | High/low, arrhythmia, technical, crisis, warning, advisory, information
Priority/levels | Yes/5. they can define the alarms level. | Cyan*, yellow**, red***, technical | Physiological levels: crisis, warning, advisory, message; technical levels: system warning, system advisory, system message | Advisory (low), serious (medium), life-threatening (high) |  | Yes
User defined | Yes | Selectable on or off, except red alarms | Yes | Yes for arrhythmia alarms | Yes | Yes
Silence active alarm | Yes | Yes | Yes | Yes | Yes | Yes
Silence duration, sec | Yes | Varies Settable ( from 60 to infinite) | Typically 60, determined by monitor | 120 | User configurable: 45/180/next alarm | User configurable
All-alarm suspend | Yes | Yes | No | Yes for telemetry | Yes | No | No
Suspend duration, sec | User configurable | 180 for telemetry, set at bedside | NA | up to 5 min | No | ? | ?
Remote settings | Complete access for monitor alarm settings | Complete access for monitor alarm settings | Yes | Yes | Yes | yes | yes
Alarms activity assessment | No | Yes: Alarm summary app | No | No | No | Yes : multi patient alarms Events | Yes : multi patient alarms Events with dedicated alarm summary equivalent
Data type | Not specified | Trend, alarm, event, wave, unit | FD, vital signs, graphic trends, limits, calipers, event directory, more | Simultaneous ECG, full disclosure (strip, hour, 24-hour, shift, selected strip, graphical trend, patient status, caliper report), event disclosure (strip shift, graphical trend, selected events), Vent Central, 12-lead rest ECG | Full disclosure, 12-lead ECG, trends (graphic and tabular), strips; | ? | ?
Export format | printer, PDF, XML for wave, | PCL5/6 printer , PDF, PNG | ? | Vary by solution | Print real time or to share file in pdf format | JPEG | JPEG
Drug | Yes | Yes, review IVPM report only, using PDF reviewer | ? | ? | ? | No | No
Hemodynamics | Yes | Yes, review IVPM report only, using PDF reviewer | ? | ? | ? | Yes | Yes
Oxygenation | Yes | Yes, review IVPM report only, using PDF reviewer | ? | ? | ? | No | No
Ventilation | Yes | Yes, review IVPM report only, using PDF reviewer | ? | ? | ? | No | No
Renal | Yes | No | ? | ? | ? | No | No
Operating system | Windows 10 (W11 expected H2 2025) Windows Server 2022 | Windows 10/Win Server 2016 | win 7 embedded | Linux | Windows | Windows XP New central station CNS 6801: Win10 | Windows 7 SP1
Database management system | Proprietary | Microsoft SQL | NA | Versant, optional RAID 1 |  | NA
Memory | Q2 2025 release | 4 GB RAM (central); 12 GB server | ? | > 1GB |  | 2 GB | 2 GB
Central Station Data Storage | RAID | SSD | SSD |  |  | ? | SSD
Communication protocols | Ethernet, TCP/IP (Broadcast or Multicast) | Ethernet, TCP/IP, Multicast (Unicast iX surv to Server) | Ethernet, TCP/IP (Broadcast only° | TCP/IP, IP Multicast |  | Ethernet, TCP/IP
Architecture | Distributed or client/server | Client/server | Peer to Peer | Client-server for telemetry Distributed for Monitor | Distributed | Star Bus Topology
Max no. of central stations on single network | Per network capacity, N | 250 | Max 15 | 256 | Varies | unlimited | unlimited
Max no. of devices on single network | 3000 | 9600 (1600*6) | 1023 max | 1024 | 1000 | 1000 | 1000
VMware supported | Y | Y | ? |  | Y | ? | ?
Layer 3 supported | Y | Y | ? |  | Y | ? | ?
Cyber security | Refer to CyberSecurity latest doc | OS patching/Nod Ident/Encryption |  |  |  | N (no OS patching, no encryption, no node identification | N (no OS patching, no encryption, no node identification
HL7 outbound (vital signs) | Yes, no central station needed, or the gateway can be included into the central station | Yes (numeric + alarms) | Yes, via CARESCAPE Gateway | Yes, via Infinity Gateway | Yes , via ICS | Yes need gateway | Yes need gateway
ECG data management system | Yes (via eGateway) | Yes ECG export for Zymed Holter + ECG order management | Yes | Yes, via Infinity Gateway. 12L ECG | Yes , via ICS | Yes need gateway same than HL7 | Yes need gateway same than HL7
HL7 inbound (ADT) | Yes via eGatway (CMS not mandatory but it provides extra features) | Yes | Yes, via CARESCAPE Gateway | Yes, via Infinity Gateway | Yes , via ICS | Yes need gateway same than HL7 | Yes need gateway same than HL7
Wave Export | Yes - Strip Export (via eGateway) | Yes : Strip export ( no gateway needed) Data Warehouse ( WSQL Data base) | Yes, near real time export Via CARESCAPE Gateway , XML format | Yes Win Access API ( C and C++ script) | ? | Yes : Strip export (no gateway needed; XML file) | ?
Note on C15: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: To Be Updated after Q2 2025 with next generation of CMS